## Slide 1
MODULE 5 Incorrect Dialyzer or Dialysate
Incorrect Dialyzer or Dialysate

## Slide 2: Preventing the Use of Incorrect Dialyzer or Dialysate Is Key to Keeping Kidney Patients Safe
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 3: Incorrect Dialyzer or Dialysate
Setting up the wrong dialyzer or dialysate for a patient is a dangerous event that can result in great harm to patients.
Feasible remedies for ensuring that initial set-up errors do not result in patient harm range from patient and professional education to procedural safeguards such as technician checklists and rules for set-up procedures.
Increasing patient involvement in dialysis care and safety issues may provide another approach to safeguarding against dialyzer solution errors.
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 4: Facts about Incorrect Dialyzer or Dialysate1
  Nearly two-thirds of professionals report that the wrong dialysate has been set-up for a patient within the past three months.
  Patients who are involved in their dialysis care are less likely to report that they have had the wrong dialysate set up for their treatment than those who are uninvolved in their care (80% vs. 46%, respectively).
  The Health and Safety Survey of patients and nephrology professionals found:
    17% patients reported problems with settings on their dialysis machine;
    3% had the wrong dialyzer set up for treatment;
    2% had the wrong dialysate set up.
  1 Health and Safety Survey to Improve Patient Safety in End Stage Renal Disease, 2007
<number>
MODULE 5 Incorrect Dialyzer or Dialysate
Notes: <number>

## Slide 5: Facts about Incorrect Dialyzer or Dialysate Solutions1
  Uninvolved patients are more likely to report that they do not know if the wrong dialysate had been set up for their treatment (48% vs 18% of involved patients).
  Components of the dialysis prescription include treatment duration, model and size of dialyzer, blood flow rate, and dialysate flow rate. It should also include reference to dialysate composition (e.g., sodium, potassium, calcium, and bicorbonate concentrations). Deviation from the prescription can result in injury to the patient or inadequate treatments.
  1 Health and Safety Survey to Improve Patient Safety in End Stage Renal Disease, 2007
<number>
MODULE 5 Incorrect Dialyzer or Dialysate
Notes: <number>

## Slide 6: Facts about Incorrect Dialyzer or Dialysate Solutions
  Dialysis machine safety monitors cannot prevent adverse events. They can only detect adverse events and mitigate the related harm.
<number>
MODULE 5 Incorrect Dialyzer or Dialysate
Notes: <number>

## Slide 7: Pharmacists on Rounding Teams Reduce Preventable Adverse Drug Events in Hospital General Medicine Units Kucukarslan SN, Peters M, Mlynarek M, Nafziger DA Arch Intern Med. 2003 (Sept); 163:2014-2018
MODULE 5 Incorrect Dialyzer or Dialysate
Errors in Dialysis Treatments
Errors in Dialysis Treatment

## Slide 8: Factors Contributing to Incorrect Dialyzer or Dialysate
Use of wrong size or type of dialyzer for patient resulting in improper clearance.
Inadvertent use of acid concentrate (the acid part of a two-part bicarbonate concentrate system) in a hemodialysis machine designed to proportion acetate concentrate.
Incompatibility - the use of mismatched concentrates in a particular bicarbonate delivery system.
Crossed or improper connections can result in problems ranging from dialysate compositions that differ from the prescribed formulation to dialysates that prove fatal. Delivery of acid concentrate alone can be fatal, depending on the proportioning system used and the conductivity and other alarm limits of the dialysis machine.
MODULE 5 Incorrect Dialyzer or Dialysate
Notes: Although the frequency of falls within individual dialysis facilities might appear rare, when extrapolated nationally an estimated 15,240 end stage renal patients fall in dialysis facilities nationwide every three months.
Factors contributing to patient falls are many and varied; however, 40 percent of patients who fell reported that dizziness was the main reason for their fall.
Tripping slipping and other environmental factors along with loss of balance and general weakness are also major contributors
Once a patient falls for the first time, a downward spiral leading to more falls can develop due to the patient becoming over-cautious. They may decrease physical activities for fear of fall, which weakens physical functioning and increases the likelihood of future falls.

## Slide 9: Mechanical Factors Leading to Dialysis Errors
Ruptured dialyzer
Clotted dialyzer
Air embolism
High conductivity
Low conductivity
Low water pressure
High venous pressure
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 10: Factors Contributing to Use of Incorrect Dialyzer or Dialysate
Staff unfamiliar with procedures
Insufficient staff to handle work load
Lack of counseling or disciplinary action when procedures are not followed
Staff discomfort with reporting medical errors
Insufficient staff training
High rate of staff turnover
Lack of effective Quality Assessment and Performance Improvement program
Lack of patient education about their care
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 11: Preventing Dialysate Errors
Check the pH and conductivity of the dialysate, as well as the conductivity, pH, and temperature alarm systems, before each dialysis treatment. Checks of proper concentrate, conductivity, and pH should be included in the pretreatment check of all components and alarm systems of the dialysis machine.
Make sure that all personnel in your unit are aware of the types of dialysate concentrates available, even if you currently use only one type. Be sure that this information is included in the orientation program for new employees. Institute an ongoing educational program to keep employees informed of developments in all areas of dialysis treatment.
Whenever possible, develop (or use if available) a system of labeling (see recommendations for manufacturers) connectors and containers that prevents or minimizes crossed connections and use of mismatched concentrates. When ordering solutions, request that such systems be incorporated (e.g., color coding, different connectors) in delivered products, not shipped separately.
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 12
Store and dispense dialysate concentrates as though they were drugs. Develop a policy, management, and storage system that will effectively control the mixing and dispensing of all concentrates.
Double-check and record concentrate formulas on the patient's record. Consider a procedure for countersigning patient and storage records.
Do not dispense concentrates from large containers into smaller ones without a "keyed" dispensing system. Whenever possible, purchase concentrates in single-treatment (2½-gallon) containers.
Always dispose of concentrates remaining from the previous treatment. Do not pour remaining concentrate into another container or use in the next treatment. Replace empty or partially full containers with full ones.
Whenever possible, standardize equipment so that only one bicarbonate concentrate system is used.
MODULE 5 Incorrect Dialyzer or Dialysate
Preventing Dialysate Errors

## Slide 13
Measure the delivered dose of HD at regular intervals to ensure adequacy.
Assign patients on re-use their own personal dialyzer with their name and birth date on it.
Verify the dialyzer volume. Dialyzers intended for reuse should have a blood compartment volume not less than 80% of the original measured volume or a urea (or ionic) clearance not less than 90% of the original measured clearance.1
1. KDOQI Clinical Practice Guidelines and Clinical Practice Recommendations. 2006 Updates.
MODULE 5 Incorrect Dialyzer or Dialysate
Preventing Dialyzer Errors
Notes: Although the frequency of falls within individual dialysis facilities might appear rare, when extrapolated nationally an estimated 15,240 end stage renal patients fall in dialysis facilities nationwide every three months.
Factors contributing to patient falls are many and varied; however, 40 percent of patients who fell reported that dizziness was the main reason for their fall.
Tripping slipping and other environmental factors along with loss of balance and general weakness are also major contributors
Once a patient falls for the first time, a downward spiral leading to more falls can develop due to the patient becoming over-cautious. They may decrease physical activities for fear of fall, which weakens physical functioning and increases the likelihood of future falls.

## Slide 14
Clearly label the dialyzer with the patient’s name. Use the dialyzer only used for the same patient.
Test the dialyzer after each use to make sure it is working properly.
Test the dialyzer after rinsing and before patient is connected to venous and arterial lines for any traces of disinfectant that may remain.
Learn the possible signs and symptoms of reactions due to reuse and monitor patients for those reactions. Monitor patients for any reactions due to reuse.
MODULE 5 Incorrect Dialyzer or Dialysate
Preventing Dialyzer Reuse Errors

## Slide 15
Turn off blood pump immediately when error is noted.
Clamp both arterial and venous needle lines with hemostats.
Clamp both arterial and venous bloodlines.
Disconnect both bloodlines from needle lines.
Discard all blood lines, and everything but the dialyzer used in association with the incorrect dialyzer.
Flush both arterial and venous needles with 10cc of Normal Saline leaving syringes attached to needles.
Re-set up machine for patient using appropriate dialyzer per proper procedure.
Document incident in CV4.
Complete incident report.
Have patients complete a HIV consent form.
Draw the following lab work: Hepatitis Panel, (both B and C) and a HIV on both patients.
NOTE: Return the dialyzer to reprocessing.
MODULE 5 Incorrect Dialyzer or Dialysate
Correcting Incorrect Dialyzer in a Reuse Setting

## Slide 16: Best Practices for Preventing Dialyzer or Dialysate Errors
Develop a culture of safety and encourage a safe environment
  Implement a system for anonymously reporting "near misses“
  Perform root cause analyses on near misses and adverse events, then follow the occurrences of these events in the facility's QAPI safety program.
Review existing policies and procedures
  Verify in writing prescribed dialyzer type is used every treatment
  Verify in writing prescribed dialysate is used every treatment
  Maintain prescribed blood flow rate throughout entire treatment or arrange for adding the time spent at a lower than prescribed blood flow rate to subsequent dialysis treatments
Develop standardized systems
Develop patient safety plan
  Establish system for reporting errors
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 17: Checklist for Preventing Use of Incorrect Dialyzer or Dialysate
  Two staff members independently check and initial the dialyzer and dialysate.
  The patient checks and initials dialyzer and dialysate. Patients, educated on the dialysis procedure, dialysate and labs (connection between their potassium level and dialysate), will be able to identify if the wrong potassium bath is being used. (Staff should provide dialysis prescription updates to patients.)
  Staff and patients are educated on the dialysis treatment/procedure, explaining each step in the process, including what each identified area on the panel means, e.g., UFR, BFR, etc.
  Staff check the blood flow rate on the machine to make sure it matches what was ordered.
  Staff check the dialysate flow rate on the machine. A decrease in dialysate flow decreases toxin removal.
  If the facility reuses the dialyzer, staff check the “fill volume”.
<number>
MODULE 5 Incorrect Dialyzer or Dialysate
Notes: <number>

## Slide 18: Developing Policies and Procedures
Review current policies and procedures to ensure they meet current recommendations for preventing errors and adverse events.
Review CMS Conditions for Coverage for ESRD Facilities.
Consider using a process analysis fishbone or other root cause analysis to examine policies and procedures.
Review examples of Quality Assessment and Performance Improvement (QAPI) projects and develop QAPI projects appropriate for your facility.
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 19: Standardized Systems
Standardized systems reduce opportunities for error.
Human errors can be caused by an individual’s short-term memory, judgment impacted by long hours, or problem solving during a stressful event.
Standardizing processes and systems minimizes the possibility for errors due to the human limitations that even the most vigilant professional can have.
Standardized systems will give staff the strength to decrease variation in patient care and, ultimately, improve safety.
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 20: System for reporting errors/adverse events
All staff—clinical, clerical, housekeeping, and maintenance—as well as patients, need specific, written directions on how to report errors or adverse events.
There should be discussions to ensure that all team members clearly know what staff is responsible for responding to errors and near misses immediately.
A reporting form should be created for documentation.
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 21: Root cause analyses
Once a continuing or evolving problem is identified through data tracking, the facility should conduct a root cause analysis to determine contributing factors, including a review of systems that were in place and determining if processes were followed.
If not, there should be focus on staff training to prevent similar adverse events.
If processes were followed but did not ensure patient safety, the Patient Safety Officer and other key staff should identify risks that could lead to further incidents and create solutions to prevent future errors that could lead to patient injury.
MODULE 5 Incorrect Dialyzer or Dialysate
Notes: <number>

## Slide 22: Root Cause Analysis – Process Analysis Fishbone
A process analysis fishbone is a diagrammatic way to examine the policies, procedures, people, and equipment involved in a process leading to an outcome. It can be used to structure a brainstorming session. It immediately sorts ideas into useful categories.
Providers may use the process analysis to develop Quality Assessment and Performance Improvement (QAPI) Projects.
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 23: Process Analysis Fishbone
The Fishbone Diagram
1. Does NOT Use
Cause-And-Effect
2. Modes Are NOT
Dependent Upon
Each Other
3. Uses
Brainstorming
Primarily
4. Allows Use of
Opinion as Fact
5. Promotes Belief that
All Causes are Within
Categories Used
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 24: Fishbone Diagram Procedure
Materials needed: flipchart or whiteboard, marking pens.
Agree on a problem statement (i.e. patient safety issue). Write it at the center right of the flipchart or whiteboard. Draw a box around it and draw a horizontal arrow running to it.
Brainstorm the major categories of causes of the problem. If this is difficult use generic headings:
    Facility policies regarding the outcome (error)
    Facility procedures for implementing policies regarding the outcome (error)
    People involved in the outcome (error)
    Materials or equipment involved in the outcome (error)
    Measurement of outcome (error)
    Environment related to outcome (error)
Write the categories of causes as branches from the main arrow.
Brainstorm all the possible causes of the problem. Ask: “Why does this happen?” As each idea is given, the facilitator writes it as a branch from the appropriate category. Causes can be written in several places if they relate to several categories.
Again ask “why does this happen?” about each cause. Write sub-causes branching off the causes. Continue to ask “Why?” and generate deeper levels of causes. Layers of branches indicate causal relationships.
When the group runs out of ideas, vote on which causes of the outcome (error) seem most likely. These causes should then be used to make system-wide changes to mitigate the outcome (error). (See continuous quality improvement methodology).
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 25: Root Cause Analysis – 5 Whys
The 5 Whys can also be used to find the root cause of a problem. It leads the team to explore systemic explanations that go beyond: event-oriented excuses; the tendency to blame individuals; and a team’s inclination to first identify factors external to the system over which they perceive they have little influence.
5 Whys strategy looks at any problem and asks: "Why?" and "What caused this problem?"
Very often, the answer to the first "why" will prompt another "why" and the answer to the second "why" will prompt another and so on; hence the name the 5 Whys strategy.
MODULE 5 Incorrect Dialyzer of Dialysate

## Slide 26
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 27: True or False:
Patients do not need to be educated about their dialysis treatment.
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 28: True or False:
Patients do not need to be educated about their dialysis treatment.
FALSE
Patients who are involved in their dialysis care are less likely to indicate that they have had the wrong dialysate set up for their treatment than those who are uninvolved in their care (80% vs. 46%, respectively). Patients should be educated on the dialysis procedure, dialysate and labs (connection between their potassium level and dialysate) and be provided with updated prescriptions.
MODULE 5 Incorrect Dialyzer or Dialysate
Notes: FALSE
Recommending assistive devices may be part of effectively addressing patient physical functioning and strength issues affecting the risk of patient falls.

## Slide 29: Which of the following should done prior to starting a patient’s dialysis treatment?
Verifying the prescribed dialyzer type is being used
Verifying the prescribed dialysate type is being used
Checking the conductivity, pH and temperature alarms
Asking the patient to verify the dialyzer and dialysate
All of the above
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 30: Which of the following should done prior to starting a patient’s dialysis treatment?
Verifying the prescribed dialyzer type is being used
Verifying the prescribed dialysate type is being used
Checking the conductivity, pH and temperature alarms
Asking the patient to verify the dialyzer and dialysate
All of the above
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 31: True or False:
Setting up the wrong dialysate for a patient is a rare event.
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 32: True or False:
Setting up the wrong dialysate for a patient is a rare event.
FALSE
Nearly two-thirds of professionals report that the wrong dialysate has been set-up for a patient within the past three months.
MODULE 5 Incorrect Dialyzer or Dialysate
Notes: FALSE
Recommending assistive devices may be part of effectively addressing patient physical functioning and strength issues affecting the risk of patient falls.

## Slide 33: Additional Resources About Preventing Use of Incorrect Dialyzer or Dialyzing Solutions
Keeping Kidney Patients Safe, Patient Safety Improvement Plan Basics
http://www.kidneypatientsafety.org/toolkit.aspx
Keeping Kidney Patients Safe, Dialysis Safety: What Patients Need to Know
http://www.kidneypatientsafety.org/pdf/toolkit/RPA_dialysis_safety.pdf
Health and Safety Survey to Improve Patient Safety in End Stage Renal Disease http://www.kidneypatientsafety.org/pdf/healthsafetysurveyreports/HSSProfessionalSurveyReport_FNL_3-21-07.pdf
MODULE 5 Incorrect Dialyzer or Dialysate

## Slide 34: Additional Resources About Preventing Use of Incorrect Dialyzer or Dialyzing Solutions
MCN Healthcare Hemodialysis Checklist
http://www.mcnhealthcare.com/downloads/samples/acutehemodialysis.pdf
Cause Analysis Tools, American Society for Quality
http://www.asq.org/learn-about-quality/cause-analysis-tools/overview/fishbone.html
HDCN Quiz: Basics of dialyzer function: Extraction ratio and clearance
http://www.hdcn.com/quizzer/adeq/er21/doorway.html
MODULE 5 Incorrect Dialyzer or Dialysate